Sectioning 101

By Jen Winther

1. Trim mounted block

a. Trapezoid shape, wider base at cutting edge

b. Parallel sides

c. Slightly wider towards mount

specimen

View from top

Cutting edge

View from side

- Setting up the microtome

  - Set desired thickness usually between 5-8micro meters

  - Knife angle should be between 4 and 5 degrees

  - Knife should be securely fastened

  - Position knife away from mount

  - Secure block in mount, make sure block edge is parallel to knife edge

  - Bring knife close to, but not too close block edge

Block mount

Block mount

Knife mount

Adjusting screws

knife

- Sectioning

  - Place pencil labeled slides on slide warmer set to 40 degrees C

  - Cover slides with clean ddH2O

  - Unlock turning mechanism and push away from you, this will advance the block set micrometers each turn

  - As ribbon comes off, manipulate with tweezers, wood sticks whatever works for you. This takes practice. Be sure not to damage knife edge with tools.

  - Place ribbons on construction paper until ready to mount, be sure to keep track of order

- Mounting ribbons on slides

  - Cut ribbons to slightly smaller than available area on slide because they will expand when they hit water

  - Use two sticks, tweezers, magic whatever works to mount ribbons. Start at one end and slowly bring down to ensure straight, non-overlapping rows

  - Allow sections to rest, once fully expanded tilt slightly and suck off excess water, can wick away with kimwipe.

  - Allow slides to dry at least 4 hours or overnight.

Ribbon

Water

Slide

Ribbon

- Some staining protocols

Saffarin Fast-green Staining

100% citrisolve 1   		5 min

100% citrisolve 2   		2 min

100% EtOH 1   			2 min

100% EtOH 2         		1 min

70% EtOH 			5 min

1 % Saffarin in 50% EtOH   	10-40 min

70% EtOH 			1 min

95% EtOH				1 min

1% Fast Green in 95% EtOH	Dip 3 times

95% EtOH				30 sec

95% EtOH				1 min

100% EtOH 2			1 min

100% EtOH 1			1 min

- let slides dry and then mount with a mixture of permount and citrisolve

Toluidine Blue Staining (or any water based stain)

100% citrisolve 1   		5 min

100% citrisolve 2   		2 min

100% EtOH 1   			2 min

100% EtOH 2         		1 min

70% EtOH 			1 min

50% EtOH				1 min

35% EtOH				1 min

H2O (or buffer)			1 min

1% Toluidine Blue Stain 		1 min (follow direction for specific stain)

H20 rinse excess stain		1 min

Let air dry or blot